## Slide 1: FEGS Spatial and Temporal Analysis
Mason G. Quick
September 2019
GLM Science Meeting
FEGS Version 2 Dataset

## Slide 2
10 km
10 km
FEGS Footprint
Array of 30 radiometers
  5 x 5 main radiometer
  777 nm OI emission;
    5 additional spectral
  1080p Digital Video Camera
  10x10 km footprint at cloud-top;
  2x2 km nominal resolution;
  100 kHz sample rate
  Optical to Optical comparison
Fly’s Eye GLM Simulator (FEGS)
Projected FEGS Viewing Angles

## Slide 3: 2017 GOES-R Validation Flight Campaign
11 GLM flights, over 10k flashes, over 100k pulses
Oceanic, continental, coastal storms
Convective initiation to Supercell storms
Coordinated ground observations
Data Quality Control Removes
  Observed during a turn (bad radiometry)
  Observed at night (amplifier undershoot issue)
  High likelihood of false detection (noisy pixels)
GLM Overall Daytime Flash DE: 61%
  Range 51 – 85%
  Lower limit daytime FDE

## Slide 4: Pulse Radiance Distributions
GLM MISS
GLM MATCH
25% of the total optical energy observed by FEGS
75% of the total optical energy observed by FEGS
GLM Pulse DE: 29%

## Slide 5: Pulse Temporal Characteristics
10-90% - matches have 44% longer risetime
50-50% - matches have 38% longer fwhm
10-10% - matches have 46% longer duration
GLM MISS
GLM MATCH

## Slide 6
How does the source size affect GLM DE?

## Slide 8: Accumulated Average Pulse Radiance Radial Profiles
Median
Match: 10.5km (350km2)
Miss: 6.6km (140km2)
Campaign GLM median group footprint = 140 km2

## Slide 9: Inner Pixel Pulses
6.7% of dataset (N=6329)
Similar percentage each flight
Low power/energy through cloud-top
Relatively short duration (60% of median)
GLM detects 3.5 %
Ground networks detect 5 %
94% classification IC
IP distribution skewed toward low amplitude negative ICs
Special cases where FEGS observes the entire radiance pattern.
Max Size: 8 x 8 km

## Slide 12: End Thank You

## Slide 13: GOES-R Validation Flight Campaign (2017)
From March to May of 2017, FEGS flew as a primary payload on the ER2 as part of the GOES-R Validation Flight Campaign. Of the 16 flights, 11 were GLM focused and targeted thunderstorms. A variety of thunderstorm phenomenology were observed including oceanic, continental, and coastal storms, convective initiation and supercell storms, during both day and night at locations across the continental United States and Atlantic Ocean. During the 45 flight hours observing thunderstorms, over 10,000 lightning flashes containing over 100,000 optical pulses were observed; a factor of 10 increase over datasets collected in previous airborne lightning campaigns. These measurements were taken in conjunction with ground based observation networks such as LMAs, radar, and long range lightning detection networks.

## Slide 14: Dataset Quality Control
QC dataset removes optical pulses there were:
Observed during a turn (bad radiometry)
Observed at night (amplifier undershoot issue)
High likelihood of false detection (noisy pixels)
Quality Detections can be further filtered by:
Simultaneous detection from airborne EFCM
Simultaneous detection with GLM or Ground Network
Distance to simultaneous detection

## Slide 15: FEGS Version 2 Dataset – QC Statistics
ID | QC Filter | Flashes | % | Pulses | % | Pulse J | Flash J
0 | All Data | 14,146 | 100 | 115,213 | 100 | 4.48E-07 | 1.50E-6
1 | Flat Wings | 13,144 | 93 | 108,249 | 94 | 4.52E-07 | 1.60E-6
2 | True | 12,540 | 89 | 113,607 | 99 | 4.60E-07 | 2.36E-6
3 | Daytime | 11,781 | 83 | 102,830 | 89 | 4.29E-07 | 1.34E-6
4 | AND(1,2,3) | 9,704 | 69 | 94,850 | 82 | 4.47E-07 | 2.23E-6
5 | GN #1 match | 11,027 | 78 | 23,949 | 21 | 9.57E-07 | 3.06E-6
6 | GN #2 match | 8,291 | 59 | 10,200 | 9 | 1.20E-06 | 3.85E-6
7 | EFCM match | 5,476 | 39 | 50,865 | 44 | 6.14E-07 | 5.01E-6
8 | OR(5,6,7) | 11,433 | 81 | 63,186 | 55 | 6.35E-07 | 2.96E-6
9 | AND(4,8) | 8388 | 59 | 49,186 | 43 | 6.12E-07 | 3.28E-6

## Slide 16: GLM Daytime Flash Detection Efficiency
22% of FEGS flashes are single pulse

## Slide 18: GLM Daytime Flash Detection Efficiency Version 2 FEGS dataset
DATE | LOCATION | STORM TYPE | DAY/NIGHT | Flash Count | QC Flash Count | QC GLM FDE
03/21 | California | Active QLCS | Daytime | 2289 | 2181 | 51
04/16 | Oklahoma | Large MCS | Night/Dawn | 390 | 61 | 85
04/18 | Huntsville | Convective Initiation | Day | 348 | 316 | 44
04/20 | Toronto | Org. Convection | Dusk | 394 | 58 | 74
04/22 | Huntsville | Tornadic QLCS | Day/Dusk | 2654 | 2423 | 73
04/27 | Huntsville | Horiz. Extensive LFR | Night | 39 | - | -
04/29 | Oklahoma | Supercell | Night | 1064 | - | -
05/08 | Colorado | Severe Hail | Day | 2454 | 2311 | 54
05/12 | Louisiana | Weak QLCS | Day | 2674 | 1714 | 65
05/14 | Atlantic/KSC | Org. Convection LFR | Day | 1234 | 640 | 64
05/17 | Oklahoma | Active QLCS | Night | 606 | - | -
 |  |  | Day | 11781 | 9704 | 61
 |  |  | Night | 2365 | - | -
Spatio-temporal comparison performed by Katrina Virts

## Slide 19: FEGS Pulse (sub-flash) Statistics
DATE | LOCATION | Pulse Count | TQCD Pulse Count | TQCD GLM PDE
03/21 | California | 29934 | 28573 | 16
04/16 | Oklahoma | 1710 | 440 | 61
04/18 | Huntsville | 2370 | 1905 | 36
04/20 | Toronto | 1879 | 470 | 49
04/22 | Huntsville | 24150 | 22269 | 37
04/27 | Huntsville | 634 | 0 | 
04/29 | Oklahoma | 5324 | 0 | 
05/08 | Colorado | 21428 | 20285 | 17
05/12 | Louisiana | 16353 | 14555 | 46
05/14 | Atlantic/KSC | 7653 | 6353 | 48
05/17 | Oklahoma | 3778 | 0 | 
 | Day | 115213 | 94850 | 29
 | Night | 9736 | - | -

## Slide 20: GLM Daytime Radiance Detection Threshold Version 2 FEGS dataset
6.4 J/m2-sr

## Slide 21: Pulse Widths – IC vs CG
IC N=19289
CG N=2761
CG median pulse rise-time and widths are 100-200 us longer than IC (~20% longer).
No identifiable difference in Peak Radiance or Radiance Energy distributions.

## Slide 23: Inner Pixel Pulses
Special case where FEGS captures the entire cloud-top radiance pattern.
Max Size: 8 x 8 km
Most underfill the 3x3 array
6.7% of dataset (N=6329)
Similar percentage each flight
Low optical power/energy through cloud-top
Relatively short duration (60% of median)

## Slide 24
Estimated Peak Current Distributions
All Pulses
Inner Pulses
Peak Current (kA)