## Slide 1
2025 GLM Science Team Meeting
September 2nd, 2025
Tewa Zepp
LMX Instrument Status

## Slide 2: Introducing: LMX
Externally, LMX shows its proven GLM heritage
Internally, LMX introduces three major technology advances
  Next generation large format all-digital CMOS focal plane
  Real time event processor with adaptive thresholding
  Refined L1b filter algorithms pull more lightning from Level 0 data
LMX keeps the best parts of GLM and adds these advances to achieve improved performance
LMX seamlessly continues lightning observations from GEO
credit: Bill Line
LMX Sensor Unit
LMX Electronics Unit

## Slide 3: LMX Performance
LMX provides 100% coverage overlap over CONUS from East and West, with expanded coverage of Alaska and polar regions.
Smaller, dimmer optical pulses are prevalent in supercells. They are challenging to detect against high cloud albedo. LMX reaches deeper into the noise than GLM.
Full-frame simulated image illustrates full disk coverage of LMX, with 1.7x more pixels than GLM sampled 1.6x faster
Improved Sensitivity
Greater Coverage
Improved Resolution
East
West
86%
requirement
86%
requirement
expanded
coverage
East
LMX will detect nearly 2x more lightning pulses than GLM

## Slide 4: LMX by the Numbers
(Current predicts, subject to change)
LMX has unprecedented sensitivity to sub-pixel lightning
Parameter | Unit | GeoXO LMX | GOES-R GLM (ref)
Lens prescription | mm EFL, f number | 134 f/1.2 | 134 f/1.2
Lens field of view | deg radius | 8.8 | 8
% of GEO disk covered | % | up to 95% (86% required) | 86%
Focal plane format | pixels (N-S x E-W) | 2044 x 1480 (windowed) | 1372 x 1300
Frame rate | fps | 774 | 503
Pixel size | µm (N-S x E-W) | 20 x 24 | 30 x 30 (variable pitch)
Ground sample distance @ nadir | km | 6.4 | 8.0
ADC bit depth | bits | 12 | 14
Downlink data rate | Mbps | 75 | 7.7
Event data structure size | bits | 96 | 64
Probability of detecting a dim pulse (6.4 – 8.8 µJ/m2/sr) of 6 km Ø | % | 58 | 10

## Slide 5: Three Innovations for LMX
The CIS326 is a bespoke high speed, large format monolithic CMOS imager with low read noise and on-chip ADC. It generates 30 Gbps of raw video data, ~3x GLM (First Light achieved June ‘24)
Initial characterization test completed
The new RTEP algorithm employs each pixel as a jitter sensor to measure its noise and adapt its individual threshold to a target TNR (Threshold to Noise Ratio)
Cluster-before-filter architecture provides spatial and temporal context for coherency calculations to pluck more lightning out of noise, including the first event in a flash
Next Gen Event Detection Algorithm
Next Gen CIS326 from Teledyne e2v
Next Gen Ground Processing Algorithms
Uniform false events
(in cyan)
Back-side illuminated CIS326
L1b product is lightning events. Additional internal meta-data:
1-minute calibrated full disk background,
Threshold image
Data Quality map
Photo: Teledyne e2v

## Slide 6: Status & Conclusions
Lockheed Martin has innovated and invested in hardware and algorithms for LMX, shaped by the lessons learned from over 8 years of GLM operational experience
LMX will provide consistent and reliable data products to end users of lightning mapper data
No data overflows
Richer data quality indications
Fewer data artifacts
More detected lightning
Program Status
  SRR in March 2025, PDR planned for May 2026
  Technical Status:
    Technology maturation for TRL6
    CMOS FPA development with Teledyne e2v
    Instrument SU & EU designs
    GPA Updates
LMX will serve end users into the 2050s
Credit: Giorgio Montersino / Flickr / CC BY-SA 2.0
Credit: C. Clark / NOAA / Public Domain
Credit: NOAA NESDIS

## Slide 7: Thank you
© 2025 Lockheed Martin Corporation. All Rights Reserved. This presentation was made possible by work performed under NASA contract 80GSFC25CA013